## Slide 1
SKILL ENHANCEMENT COURSE (SEC) PROJECT
Design and Development of an Intelligent USB-Based Electronics Platform
Eight bench instruments on one USB cable, with Proto, an AI engineer that plans every test and explains every measurement.
PROJECT TEAM
Avi Prakash Jaiswal
24294917
Soham Pandit Nalawade
24294917002
Harshit Saini
24294917086
Anshul
24294917037
SUPERVISOR
Prof. G.S. Chilana
ONE DEVICE, EIGHT INSTRUMENTS
Oscilloscope
Waveform generator
Logic analyser
Pattern generator
Protocol decoder
Power supply 1
Power supply 2
Multimeter
+ Proto
an AI engineer that runs the tests and gives a verdict in plain words
Notes: Speaker: Avi · about 0:45 Good morning. We are presenting our SEC project, Design and Development of an Intelligent USB-Based Electronics Platform, under the supervision of Prof. G.S. Chilana. I am Avi, with Harshit and Anshul; our team member Soham Pandit Nalawade could not be present today. In one line, our project is one small device that does the work of eight lab instruments, connects to a laptop with one USB cable, and has an AI engineer, called Proto, that runs tests and explains the results in plain language.

## Slide 2
MOTIVATION AND OBJECTIVES
Why one intelligent platform instead of a bench of instruments
Intelligent USB-Based Electronics Platform · SEC Project
2 / 7
THE PROBLEM
Costly and bulky
A complete electronics bench needs five to eight separate instruments.
Disconnected
Each instrument has its own interface; results are hard to combine and record.
Hard to interpret
Students can read a number but often not what it means, or what to fix.
Scope
Wavegen
Logic
Pattern
Decoder
PSU 1
PSU 2
DMM
1 USB
device
OBJECTIVES
1
Integrate eight instruments into one compact, USB-connected unit
2
Design an FPGA-based acquisition and signal-generation core
3
Develop a single software workbench for every instrument
4
Embed Proto an AI engineer that plans tests and explains results in plain language
5
Keep it affordable for student labs, with a low-cost production design
Notes: Speaker: Avi · about 1:00 Why did we build this? A normal electronics bench needs five to eight separate instruments: an oscilloscope, a function generator, a logic analyser, a multimeter and power supplies. They are costly and bulky, each has its own controls, and students often see a number without knowing what it means or what to fix. So our objectives are: put all eight instruments in one USB unit; build the core on an FPGA; give everything one software screen; add Proto, an AI engineer that plans tests and explains results; and keep it affordable for student labs.

## Slide 3
HARDWARE DESIGN
System architecture: one FPGA core, one USB link
Intelligent USB-Based Electronics Platform · SEC Project
3 / 7
CH1 · CH2 in
Analog front end
AD8066 buffers, relay range switching
2× ADC
AD9226 · 12-bit · 65 MS/s
W1 · W2 out
Output stage
buffer and level shift
Dual DAC
AN9767 · 14-bit · 125 MS/s
Logic 0–15
Level shifting
74LVC245 · 16 channels
Direct I/O
logic analyser + pattern generator
DMM in
Precision input
REF5025 reference
24-bit ADC
ADS1256 multimeter
V1 · V2 out
Programmable supplies
XL6009 regulators
Setpoint + sense
MCP4728 DAC · INA226 monitor
FPGA core
Tang Primer 20K (Gowin)
Trigger engine
Capture buffers
DDS signal synthesis
Pattern sequencer
Protocol decode (I²C · SPI · UART)
USB packetiser
USB 2.0
High-Speed
Host computer
Device driver + API
Workbench software
Data logging
Proto
AI engineer: plans tests, reads results, explains
Notes: Speaker: Harshit · about 1:15 This is how the hardware is organised. Read it from left to right. Each port on the left goes through a front end that protects and scales the signal, then through a converter. The two oscilloscope inputs use AD9226 chips, 12-bit converters that take 65 million readings per second. The waveform outputs use a 14-bit dual DAC. The 16 logic lines pass through a 74LVC245 buffer, the multimeter uses a precise 24-bit ADC, and the two power supplies are set by a small DAC and measured by an INA226 current monitor. Everything meets in the FPGA, the brain of the device. It handles triggering, captures data, generates waveforms and decodes buses, then sends everything over USB 2.0 to the computer, where our software and Proto run.

## Slide 4
HARDWARE DESIGN
Instrument interface: the front panel and what each port does
Intelligent USB-Based Electronics Platform · SEC Project
4 / 7
Instrument | Port | Key specification | Hardware
Oscilloscope | CH1, CH2 | 2 channels · 12-bit · 65 MS/s | AD8066 front end + 2× AD9226
Waveform generator | W1, W2 | 2 channels · 14-bit · up to 125 MS/s | AN9767 dual DAC, FPGA DDS
Logic analyser | Logic 0–15 | 16 channels | 74LVC245 + FPGA capture
Pattern generator | Logic 0–15 | 16 channels, programmable sequences | FPGA sequencer
Protocol decoder | Logic 0–15 | I²C · SPI · UART | FPGA + host software
Programmable supplies | V1, V2, COM | 2 outputs, voltage and current monitored | MCP4728 + XL6009 + INA226
Multimeter | DMM, COM | 24-bit resolution | ADS1256 + REF5025
Notes: Speaker: Harshit · about 1:00 This is the front panel the user actually touches. CH1 and CH2 are the oscilloscope inputs. W1 and W2 are the signal generator outputs. The 16-pin header is for logic signals: the logic analyser, the pattern generator and the protocol decoder all use it. V1, V2 and COM are the two programmable power supplies, DMM is the multimeter input, and USB-C connects to the laptop. The table shows each instrument, its port, its key specification and the chip that does the job.

## Slide 5
SOFTWARE DESIGN
Software interface: every instrument and Proto in one workbench
Intelligent USB-Based Electronics Platform · SEC Project
5 / 7
SOFTWARE STACK
Proto, AI engineer
plans tests, explains results
Workbench UI
scope, logic, decoder, supplies, DMM views
Device API
commands, streaming, logging
USB transport
USB 2.0 high-speed data link
FPGA gateware
trigger, capture, DDS, decode
Top: what the user sees. Bottom: what runs on the hardware.
Notes: Speaker: Anshul · about 1:00 This is our software design. On the left you choose an instrument. In the middle is the oscilloscope screen with two waveforms, and below it the logic view decoding an I2C message into readable bytes. On the right are automatic measurements like frequency, voltage and rise time, and the Proto panel, where you can ask a question in plain English. On the right side of the slide is the software stack, from bottom to top: the FPGA logic, the USB link, the device interface, the workbench screen, and Proto at the top.

## Slide 6
INTELLIGENCE AND RESULTS
Proto, the AI engineer: from a measurement to a verdict
Intelligent USB-Based Electronics Platform · SEC Project
6 / 7
HOW PROTO RUNS A TEST
1
Connect the board
2
Power up, watch the current
3
Check every supply rail
4
Drive the inputs
5
Capture the outputs
6
Decode the buses
7
Verdict in plain words
8
Suggest the fix
Frequency response of an RC filter
Wavegen sweeps, scope measures, Proto compares with theory: f_c = 1/(2πRC) ≈ 1.59 kHz.
Supply-rail check with tolerance
DMM and supply monitors read each rail; anything outside ±5 % is flagged.
PROTO'S VERDICT
“The 3V3 and 5V rails are fine, but the sensor supply is 2.71 V, 18 % low. Check the regulator enable pin and the 10 µF output capacitor, then run the test again.”
Graphs show illustrative data for the test method.
Notes: Speaker: Anshul · about 1:15 Proto is what makes the platform intelligent. It follows a fixed test flow: connect the board, power it up while watching the current, check every supply rail, drive the inputs, capture the outputs, decode the buses, give a verdict and suggest the fix. The first graph shows a frequency-response test of an RC filter. The generator sweeps the frequency, the scope measures the output, and Proto compares it with theory: the cutoff is one over two pi R C, about 1.59 kilohertz. The second graph checks each supply rail against a five percent tolerance. Here the sensor supply is 2.71 volts instead of 3.3, so Proto flags it and suggests what to check. These graphs show the test method with example data.

## Slide 7
PLAN AND CONCLUSION
Methodology, outcomes and future scope
Intelligent USB-Based Electronics Platform · SEC Project
7 / 7
DEVELOPMENT METHODOLOGY
1
Requirements and architecture
2
Component selection and BOM
3
Schematic and PCB design
4
FPGA gateware
5
Workbench software
6
Proto integration
7
Testing and validation
EXPECTED OUTCOMES
One USB unit replaces eight bench instruments
Tests run as plans, not as manual steps
Measurements explained in plain language by Proto
Affordable enough for every student lab bench
FUTURE SCOPE
A probe fixture that reaches every net of a board automatically
More protocol decoders (CAN, I²S, 1-Wire)
Remote lab access over the network
Proto-generated lab reports from recorded measurements
Thank you
Notes: Speaker: Avi · about 1:00 Our method has seven phases: requirements and architecture, component selection, schematic and PCB design, FPGA logic, the workbench software, Proto integration, and finally testing and validation. The expected outcome is one USB unit that replaces eight bench instruments, runs tests as plans, and explains results in simple language at a cost student labs can afford. In future we want to add a probe fixture that reaches every point on a board automatically, more protocol decoders, remote lab access and automatic lab reports. Thank you. We are happy to take your questions.